Development and Validation of a Novel Posture Monitoring System(v.2.4-Sai’s part: improvements are done in calibration routine and feature selection process- added references in the back and matched it with the introduction paragraphs-place all pic at end of doc until final draft-included some image formatting and some equ involving- involves changes in grammar in everything below intro and addition of wrap fun formula, thethaAccx, and explanation for form, and patients’ data)

Sai Sri Harini.C
Biomedical Department
PSG College of Technology
Coimbatore, India

Syed Althamish ur Rahman
Biomedical Department
PSG College of Technology
Coimbatore, India

Sarvesh Rajavel.R
Biomedical Department
PSG College of Technology
Coimbatore, India
Kefira Jane.J
Biomedical Department
PSG College of Technology
Coimbatore, India

Abstract—In contemporary society, posture-related musculoskeletal disorders are on the rise due to extended periods of sedentary behavior and repetitive tasks with people spending most of their time slouching and hunching over desktop screens, tablets and mobile phones. This continuous behavior over extended periods of time causes severe upper and lower backpain that subsequently leads to cervical spondylosis and bulging of lumbar discs. In response to the prevalence of this issue, we have developed a posture monitoring system realized through a shoulder strap to capture the siting posture of individuals. This is achieved by utilizing 3 IMU sensors to monitor and Machine Learning Algorithms to classify. The classification algorithms were trained with data acquired from 12 healthy individuals and the bests chosen one is the KNN algorithm with an accuracy of 99%. The system was further validated with a real time validation routine and was identified with 80% accuracy.
Keywords—IMU, Spinal Posture, Data Classification
# Introduction
Sitting is one of the most common postures in daily life [1,3,5,6], with studies indicating that people spend around six hours sitting each day. Back pain is a prevalent and concerning condition, recognized as a leading cause of disability globally and a frequent reason for medical consultations (18,19). Research indicates a notable growth in the occurrence of chronic low back pain over the years, with one study documenting an increase from 3.9% to 10.2% between 1992 and 2006 (20). . The occurences of low back pain in adults is reported to be 70-80%, mostly caused as a result of faulty postures and sedentary lifestyles (17). Chronic back pain develops gradually due to sedentary behavior, often exacerbated by prolonged use of digital devices.
The most frequent activity while sitting is using mobile phones, with an estimated 4.77 billion users worldwide as of 2017. This number was projected to increase to 5.07 billion by 2019 [7]. As more essential services are integrated into smartphones, and with the growing number of mobile applications, this has further promoted a sedentary lifestyle. Alarmingly, a recent study reveals that looking down at a phone is comparable to placing 60 pounds of pressure on the neck [11]. Correspondingly, another study highlighted that a condition known as "Text Neck," can develop into an extreme condition of Kyphosis [8]This disease is an excessive spine curve because of abnormal rounding of upper back. which can lead to breathing problems that may eventually require a surgical treatment. However, a majority of Kyphosis cases can be prevented and cured with exercises that improve sitting posture[12].
Chronic Low Back Pain (CLBP) is another significant cause of disability, affecting similar populations across various countries and cultures, with one in five adults experiencing CLBP each year [1]. In industrialized nations, chronic pain is rapidly becoming the foremost health issue, contributing to annual low back pain costs of $100-$200 billion [3].
Acute back pain often results in considerable discomfort and restricted mobility, as highlighted by multiple studies. Von Korff and Saunders[21] reported that 66-75% of primary care patients still experienced mild back pain one month after treatment, with 20-25% facing significant activity limitations. Additionally, Makris et al. [22] explored the connection between disabling back pain and mobility impairments in older adults. Overall, back pain significantly affects an individual’s quality of life, reducing productivity, limiting movement, and causing emotional distress. Studies have found that slouched or hunched posture not only reduces energy levels but also negatively impacts mental well-being, such as happiness and depression [9]. Poor posture also affects workplace productivity, with approximately 75% to 85% of worker absenteeism being attributed to recurrent or chronic back pain [5].
Despite its prevalence, simple preventive measures can effectively mitigate these disorders. According to medical research, most cases of severe Kyphosis can be treated and prevented through exercises aimed at improving posture and maintaining a straight spine [12]. A study by Robertson et al. [13] found that musculoskeletal risk decreased after 16 months of ergonomic posture training for seated individuals. Further studies by Choobineh et al. [14] and Menendez et al. [15] demonstrated that ergonomic interventions could reduce musculoskeletal discomfort and related symptoms. Additionally, research by Taieb-Maimon et al. [16] showed that posture risk diminished after three weeks of an experiment using a camera to display the seated individual’s sagittal posture. Therefore, maintaining proper spinal posture is possible and essential, which is the goal of our system’s design.
Several systems have been developed to address this issue, comprising of three main components as defined by Ferdews Tlili et al. [10] in his review of posture monitoring systems. The individual components are data collection, data analysis, and feedback . Our system incorporates all these elements, with a particular focus on analyzing the collected data while in a seated position.
# System Architecture
Our system's operation primarily depends on the data acquired by IMU sensors placed on the body, which are then pre-processed to extract features. These features are then fed to a classifier to identify the user's sitting posture, therefore notifying the user.
## Design:
Our system is designed to extract movement and orientation-related data during sitting posture from three different locations. These locations are:
- The Acromion Process
- The T5 vertebra
- The Sternum (SAII)
The Acromion Process and T5 vertebra were chosen due to their remarkable accuracy in existing posture monitoring systems that use IMU sensors. This is understandable due to the significant deviation in location and orientation during postural changes. The thoracic junction is used to counterbalance the system's bias toward the posterior of the body with the T5 vertebra. This location proved to be the best for extracting data from the anterior of the body.
The placement of these sensors was achieved using a Neoprene Velcro shoulder strap, as described in Figure 1. This adjustable shoulder strap covers all three mentioned areas and was used to embed the sensors within. Care was taken to avoid letting the shoulder strap's influence affect the user's normal posture. The Acromion and T5 sensors were fixed, whereas the chest sensor was adjustable to relocate it to the correct position for individuals with different body widths.

The sensors used for this purpose were MPU-6050s, which are capable of acquiring 6 degrees of freedom (DoF), namely, acceleration and rotational speed on all three axes. This data could be transmitted through the I2C medium, allowing fast multi-device communication, aiding our purpose, and operating at low power.
The processing unit was located near the sensors and was embedded within the strap near the Acromion Process, as this area had enough space to hold the MPU. The processing unit was connected to the sensors via a series of wires designed to align with the strap's structure. Some limitations were noted in its placement, with reports of discomfort due to its inflexible structure. The objective of our system is to improve processing and classification; therefore, the processing unit's size was reduced as much as possible but ideally needs to be flat and flexible with zero protruding height. The Raspberry Pi Pico was chosen for its small size and ample memory (20kb SRAM). It also includes built-in libraries for the future prospect of embedding ML algorithms. Additionally, it is low power.
A BLE module was used to transfer relevant data to the classifier on the PC. The HC-05 module was selected for its accessibility to high speeds and low power operation.
The classifier was run in the background on the user’s PC. It classified data received from the BLE module and presented relevant information to the user as needed. The classifier chosen for this purpose was KNN, as it demonstrated better performance in a set of experiments detailed in the following sections. The classifier took a mean time of 1.5 seconds for inference.
## Operation
Once the shoulder strap embedded with the sensors and processing unit is worn by the user, they are expected to perform a calibration routine. This routine is conducted to eliminate the zero-error arising from the gyroscopic measurements for each user. During this 10-second routine, the user sits in a static neutral posture, as depicted in Figure 1. The zero-error bias is calculated by averaging the static gyroscopic signal.
{gzBias} rsub {sa} = {1} over {N} sum from {i = 1} to {N} {{g} rsub {sa} left [i right ]}
Each sensor outputs a 6-DoF measurement, namely, acceleration and rotational speed on all three axes. Each sensor is sampled at 45Hz, given the processing limits of the Pico. While the acceleration signal is not pre-processed, the gyroscopic signals are altered by subtracting the zero-error bias value calculated during the calibration routine. This adjustment allows the gyroscope to output true zero values. The processing unit then acquires these signals and processes them to extract new information, known as tilt. The tilt is calculated in two ways: using a complementary filter and using only acceleration signals.
The complementary filter considers both acceleration and gyroscopic observations to calculate the tilt using the formula depicted below. The complementary filter effectively combines the low-frequency stability of the accelerometer (which measures tilt based on gravity but is susceptible to noise from linear movements) with the high-frequency responsiveness of the gyroscope (which tracks rotation but drifts over time). This fusion helps maintain a more reliable and stable estimate of tilt angles, reducing noise and drift. It

applies a weighted combination of the accelerometer and gyroscope data.

The optimal weight chosen is 0.95. The formula is further advanced to wrap the signal around +180 to -180 degrees to better compensate for the drift accumulated by gyroscopic readings. Despite the complementary filter's advantage over the acceleration-based tilt, due to the static nature of sitting postures, both methods are taken into account.

θX​[n] is the filtered tilt angle for the X-axis at the current loop iteration.
θX[n−1]\theta_X[n-1]θX​[n−1] is the tilt angle from the previous iteration.

- AccX\theta_{\text{AccX}}θAccX​ is the angle derived from accelerometer data for the X-axis, calculated using
- θ˙GyroX​⋅Δt is the change in angle from the gyroscope data between sampling period, where θ˙GyroX​ is the angular velocity from the gyroscope for the X-axis, and Δt is the time step between each sample.
α is the complementary filter weight which we’ve found best operated at a value of 0.95

{θ} rsub {x} left [n right ] = wrap left (α left ({θ} rsub {AccX} + wrap left ({θ} rsub {x} left [n − 1 right ] + {dot {θ}} rsub {GyroX} × ∆ t − {θ} rsub {AccX} , 180 right ) right ) + left (1 − α right ) × {θ} rsub {AccX} , 180 right )
Where wrap is defined as above, where:
x is the input angle.
L is the limit (e.g., L=180Deg)
⌊y⌋ is the floor function, which returns the greatest integer less than or equal to y.

Where thethaaccx is the x axis tilt derived from acceleration data of a sensor [ref?]
From all the extracted features, the meaningful ones are sent to the classifier via the BLE module using a UART module to communicate with the processor. The classifier then identifies the user's posture and displays it to them via an application.{saii}
# Feature Selection and Training{saii}
For the purpose of analyzing the data and training the ML algorithms with respect to it, appropriate data needs to be acquired. The following section highlights how that data was acquired.
## Data set Acquisition:{saii}
The data was acquired from 12 healthy individuals with a mean age, height, and weight of 23y, 166cm, and 70Kg, respectively. Among 12 individuals 8 were males and 4 were females. Their respective data is presented in Table 1.
The data acquisition routine began with the user undergoing the calibration routine outlined above. They were then asked to relax and begin with the first posture, which was neutral. Each user was seated on a stool with no back support and was requested to remain in a specific posture for 10-15 seconds. The user was asked to replicate the postures depicted in Figure 4. The order of the postures is as follows:
- Neutral
- Mild Slouch
- Moderate Slouch
- Extended Slouch
- Mild Hunch
- Extended Hunch
- Right Hunch
- Left Hunch
- Lean Right
- Lean Left
Approximately 94,000 data points were recorded, with about 9,000 data points per posture.
				Mean	Variance
Age	20	21	21	21	21	20	47	18	22	20	19	24	23	60
Height / (cm)	180	164	169	170	168	165	157	157	167	159	165	165	166	40
Weight(kg)	80	57	86	75	85	64	82	56	64	51	70	72	70	141

## Feature Set Selection:{saii}
Including the raw data measurements from each sensor and the tilt angles derived using two different methods, there is a total of 33 features. Each sensor provides 11 features: 6 DoF raw data measurements and 5 tilt angles. The X and Y axes are derived from both complementary filter and acceleration data, while the Z axis is derived only from acceleration data. Some features were deemed redundant, as detailed in Table 5.
The ideal posture monitoring system would be a stand-alone, wearable system that does not require external processing. To support this goal, a feature reduction/selection process was carried out to determine the most relevant features. The method used to achieve feature selection and extraction is principal Component Analysis (PCA), forward selection, and backward selection, all performed using k-fold cross-validationare discussed below. Details about the techniques are described below:
#### Principal Component Analysis (PCA): PCA is a widely-used dimensionality reduction technique that transforms the original set of features into a smaller set of uncorrelated components, called principal components. In this case, PCA was applied to the 33 features collected from the MPU6050 sensor, which included accelerometer, gyroscope, and tilt data. PCA helped identify which combinations of features were most critical for distinguishing between the 10 posture labels (Neutral, Slouching, Hunching, Leaning).

### Forward Selection: Forward selection is a stepwise feature selection method where the model starts with no features, and features are added iteratively based on their performance in improving the model’s predictive ability.

### Backward Selection: In contrast to forward selection, backward selection begins with all 33 features included in the model. Features were removed iteratively based on their contribution to the model’s performance.

### K-Fold Cross-Validation: To ensure that the feature selection methods were robust and did not result in overfitting, k-fold cross-validation was applied throughout the process. In k-fold cross-validation: • The dataset was split into k subsets (folds). • Where the value of k is taken as 5. • The model was trained on k-1 folds and tested on the remaining fold. This process was repeated k times, with each fold serving as the test set once. • The average performance across all folds was calculated to evaluate the model.
By using k-fold cross-validation during PCA, forward selection, and backward selection, the most relevant features were chosen based on how well they generalized across different parts of the dataset.

After applying PCA, forward, and backward selection methods, the most relevant features were selected from the original 33. These selected features were those that contributed most to distinguishing between the 10 posture labels. This selection process not only improved the model’s accuracy but also reduced its complexity, making it more efficient for training and prediction. After applying the feature selection methods—PCA, forward selection, and backward selection with k-fold cross-validation—the final selected features were the complementary tilt values of only the two sensors placed at the acromion process and at T5.
These features were found to be the most relevant for distinguishing between the 10 posture labels: Neutral, Slouching (Mild, Moderate, Extensive), Hunching (Symmetric-Moderate, Symmetric-Extensive, Right, Left), and Leaning (Right, Left). This selection streamlined the dataset, reducing it from the original 33 features to just the 6 complementary tilt values, making the model more interpretable and faster to train without sacrificing performance.

-
-
-

## Classifier Selection:{saii}
Multiple ML models were tested to find the best-performing one. The KNN model, which used only 2 or 3 user-selected features (complementary tilt angles in the X and Y axes), provided 94% accuracy. The performance details of each classifier are shown in Table 9. The training and testing datasets were divided by a 0.75 ratio.
# Result{saii}
Our system is designed to efficiently and accurately monitor sitting posture. The final design comprises three sensors whose raw data is used to extract a derived feature known as tilt. The complementary filter's tilt derivation was found to be the best-performing variation. The KNN classifier, which provided 94% accuracy, was deemed the best model for identifying posture. Although the classifier was trained with a limited dataset of only 12 people, the system would benefit from more data. It was also found that an additional third sensor on the posterior of the body reduced the classifier's accuracy, rendering it unnecessary.
# Conclusion:{Saii}
Posture monitoring with IMU sensors is a challenging but achievable task, with our system providing 94% accuracy. This system is ideal for monitoring sitting posture during seated work. IMU sensors provide maximum privacy while maintaining decent accuracy. Due to its low-power nature, the

system can be used anywhere, acquiring postural data over long periods with minimal effort from the user. However, the design could benefit from improvements, such as reducing the size of the processing unit and increasing its flexibility through the use of flexible electronics. The system's performance with the general population could improve with access to a larger and more diverse dataset, including individuals of varying ages, weights, and heights. Furthermore, the calibration routine at the start of each monitoring period could be replaced with more efficient, less time-consuming methods that do not impose restrictions on the user. The shoulder strap could be replaced with textile-embedded electronics to enhance user comfort. Finally, incorporating the ML algorithm into the processing unit would allow the system to function as a stand-alone device.

##### References
1. El-Sayed B, Farra N, Moacdieh N, Hajj H, Haidar R. A novel mobile wireless sensing system for real-time monitoring of posture and spine stress. *IEEE Trans Neural Syst Rehabil Eng* 2011; 13: 18-23.
2. [Brigante CMN, Abbate N, Basile A. Towards Miniaturization of a MEMS-Based Wearable Motion Capture System. *IEEE Transact Industrial Electronics* 2011; 58: 246-278.
3. Nemec D, Janota A, Hrubos M, Simak V. Intelligent Real-Time MEMS Sensor Fusion and Calibration. *IEEE Sensors J* 2016; 16: 1558-1748.
4. Fachin F, Nikles SA, Dugundji JA, Wardle BL. Analytical extraction of residual stresses and gradients in MEMS structures with application to CMOS-layered materials. *J Micromech Microeng* 2011; 21: 095017.
5. Donatell GJ, Meister DW, O’Brien JR, Thurlow JS, Webster JG, Salvi FJ. A simple device to monitor flexion and lateral bending of the lumbar spine. *IEEE Transact Neural Syst Rehab Eng* 2005; 13: 1534-4320.
6. Garcia E, Ding H, Sarela A, Karunanithi M. Can a mobile phone be used as a pedometer in an outpatient cardiac rehabilitation program. *IEEE/ICME Int Confer Complex Med Eng (CME)* 2010.
7. Statista (2017) Number of Mobile Phone Users Worldwide 2013-2019. Statista. https://www.statista.com/statistics/274774/forecast-of-mobile-phone-users-worldwide/
8. Jung, S., Lee, N., Kang, K., Kim, K. and Lee, D. (2016) The Effect of Smartphone Usage Time on Posture and Respiratory Function. *Journal of Physical Therapy Science,* 28, 186-189.
9. Riskind, J.H. and Gotay, C.C. (1982) Physical Posture: Could It Have Regulatory or Feedback Effects on Motivation and Emotion? *Motivation and Emotion,* 6, 273-298. https://doi.org/10.1007/BF00992249
10. F. Tlili, R. Haddad, Y. Ouakrim, R. Bouallegue and N. Mezghani, "A Review on posture monitoring systems," 2018 International Conference on Smart Communications and Networking (SmartNets), Yasmine Hammamet, Tunisia, 2018, pp. 1-6, doi: 10.1109/SMARTNETS.2018.8707392.
11. OLGA KHAZAN, “What Texting Does to the Spine,” The Atlantic, November 2014.
12. American Academy of Orthopedic Surgeons, “Kyphosis (Roundback) of the Spine“, August 2016.
13. Robertson, M.; Amick, B.C.; DeRango, K.; Rooney, T.; Bazzani, L.; Harrist, R.; Moore, A. The effects of an office ergonomics training and chair intervention on worker knowledge, behavior and musculoskeletal risk. *Appl. Ergon.* 2009, **40**, 124–135.
14. Choobineh, A.; Motamedzade, M.; Kazemi, M.; Moghimbeigi, A.; Heidari Pahlavian, A. The impact of ergonomics intervention on psychosocial factors and musculoskeletal symptoms among office workers. *Int. J. Ind. Ergon.* 2011, **41**, 671–676.
15. Goossens, R.H.M.; Netten, M.P.; Van Der Doelen, B. An office chair to influence the sitting behavior of office workers. *Work* 2012, **41**, 2086–2088.
16. Taieb-Maimon, M.; Cwikel, J.; Shapira, B.; Orenstein, I. The effectiveness of a training method using self-modeling webcam photos for reducing musculoskeletal risk among office workers using computers. *Appl. Ergon.* 2012, **43**, 376–385.
17. Cădar, I. D. and Laura Pop. “Correlations between cervical spine posture and low back pain.” *HVM Bioflux* 7 (2015): 178-181.
18. Traeger, Adrian C. et al. “Diagnosis and management of low-back pain in primary care.” *Canadian Medical Association Journal* 189 (2017): E1386 - E1395.
19. Lee, Mj. “Epidemiology of back pain.” (2012).
20. Freburger, Janet K. et al. “The rising prevalence of chronic low back pain.” *Archives of Internal Medicine* 169 3 (2009): 251-8.
21. von Korff, Michael and Kathleen Saunders. “The Course of Back Pain in Primary Care.” *Spine* 21 (1996): 2833–2837.
22. Makris, Una E. et al. “Restricting Back Pain and Subsequent Mobility Disability in Community‐Living Older Persons.” *Journal of the American Geriatrics Society* 62 (2014): 2142 - 2147.